## Slide 1: Displays & Image Processing
Chapter 15

## Slide 2: Displays
Display refers to the screen on which your images are shown.
In diagnostic ultrasound this is most often a cathode ray tube, or a flat panel monitor.

## Slide 3: Bistable vs. Gray Scale
There are two choices that we can have in image display.
  Bistable
    White or black, on or off
    high contrast, narrow dynamic range, poor contrast resolution

## Slide 5: Bistable vs. Gray Scale
Gray Scale
  many levels of brightness. (white light gray, medium gray, dark gray etc)
    Assignment of a different shade of gray for each echo amplitude
    Different colors represent different signal strengths
    Low contrast, wide dynamic range, improved contrast resolution

## Slide 6: Controls of a display (Just like the controls on your TV)
Contrast
Determines range of brilliancies that are displayed.
Bistable images are high contrast.
Brightness
Determines the brilliance of the signals displayed

## Slide 7: Contrast Resolution
Shades of gray in an image
Good contrast beagle!
Poor contrast beagle!

## Slide 8: Brightness
Bright!
Dark!

## Slide 9: Scan Converters
Changes the data format from US signals to that of a television
Scan converters make gray scale displays possible by storing the image information and then displaying it on a Cathode Ray tube
  Storage is called "writing"
  Later, image data is "read" for display on the monitor
This also allows the image to be manipulated/altered between storing and displaying.
  Ex. Black and white inversion

## Slide 11: Analog and Digital
Analog
  real world
    a variable attains a continuum of values
      Actual weight of a beagle
      True length of a beagle
Digital
  computer world
    a variable attains only discrete values.
      Measuring the Beagle’s weight on a digital scale
Oh dear…

## Slide 12: Analog Scan Converter
Originally made gray scale display possible
Funnel shaped tube with an electron gun
Divides pictures into a dielectric matrix with an electrical storage element (silicon wafer)
Electrons from the gun strike these elements and the charge is stored
Stored charges are later read to retrieve information

## Slide 13: Analog Scan Converter
Advantages of the Analog Scan Converter
    Excellent spatial resolution/image detail
Disadvantages
    Image fade-charges on silicon wafer dissipate
    Image flicker-constant switching between read and write modes
    Drift-inconsistent pictures from day to day
    Deterioration-tube ages and image degrades

## Slide 14: Digital Scan Converter
A computer memory and microprocessor used to digitize images.
  Converts image data into numbers which are stored in memory
  Numbers are processed and re-translated for display as an image.
RAM - computer memory (random access memory)

## Slide 15: Digital Scan Converters - Advantages
Uniformity – consistent gray scale quality throughout the image
Stability – does not fade or drift
Durability – not affected by age or heavy use
Speed – nearly instant processing
Accuracy – error free

## Slide 16: Pixel
Smallest element of a digital picture
Pixel density
  the more pixels per inch, the greater detail in the image
  The detail is called spatial resolution
Each pixel can only display one color at a time
The best resolution is obtained with more pixels
Spatial resolution on a digital display is determined by pixel density

## Slide 18: Bit
Derived from the term "binary digit"
Smallest amount of digital storage
Bistable - has value of 0 or 1
A group of bits is assigned to each pixel to store the gray scale color assigned to that pixel
The more bits per pixel, the more shades of gray, and the better then contrast resolution

## Slide 19: Binary Number
  Computer language
  A group of bits
    A series of 0’s and 1’s
    Example 101010101001
  Byte
    Group of 8 bits
    Example 00001011

## Slide 20: Calculating the number of Gray Shades that a collection of bits can represent
1. Find out how many bits are assigned to each pixel
2. Multiply the number 2 by itself the same number of times as there are bits
  2^n
  n= the number of bits

## Slide 21: Example #1
  Each pixel has 5 bits
  25 or 2x2x2x2x2=32
  The largest number of shades represented by 5 bits is 32

## Slide 22: Example #2
Each pixel has 8 bits
28 or 2x2x2x2x2x2x2x2= 256
256 gray shades can be represented by 8 bits

## Slide 23
We must be able to figure out how many bits it takes to represent a certain number of gray shades.
Find the minimum number of bits that can be used by figuring out how many gray shades can be represented

## Slide 24: Example #3
How many bits are needed to represent 10 shades of gray?
    4 bits – why?
    4 bits = 2x2x2x2 = 16 shades of gray
    We cant use 3 bits because 3 bits can only represent 8 shades of gray and we need at least 10!

## Slide 25: Additional Examples
How many bits are needed to represent 11 shades of gray
    4 bits = 16 shades of gray
How many bits are needed to represent 15 shades of gray?
    4 bits = 16 shades of gray
How many bits are needed to represent 17 shades of gray?
    5 bits = 32 shades of gray

## Slide 26: Analog to Digital Conversion
Analog
  everyday, real world
Digital
  computer world
Analog to digital conversion
  Electrical signals from the transducer are analog
  Digital scan converter can only process computer information
  Analog signal must be converted into digital form for input into the scan converter
  This is performed by the analog to digital converter

## Slide 27: Digital to Analog Conversion
Information in the scan converter is digital
  Televisions are analog
  Signal must be reconverted to analog form before display on a TV
    This is done by the digital to analog converter
Analog signal from transduceranalog to digital converterdigital scan converterdigital to analog converteranalog signal display on TV (CRT)

## Slide 28: Image Processing
Pre-Processing
  manipulating the data before storage in the scan converter.
  Once preprocessing occurs data is altered forever.
  Cannot be undone or reversed
Post-Processing
  Manipulation of data after storage in the scan converter
  Changes can be reversed, restores the initial values of the image data

## Slide 29: Pre-Processing
Types of pre-processing
  Log Compression
    Lowers the high-level echoes and boosts the low-level echoes.
    Signal’s dynamic range is reduced.
    TGC changes
  Write magnification
  Fill in interpolation
  Persistence
  Spatial compounding
  Edge enhancement

## Slide 30: Post-Processing
Post- processing
  Any change after the freeze frame
  Types: black white inversion, read magnification
  Contrast variation
  3-D Rendering

## Slide 31: Pre and Post Processing
Preprocessing is any image manipulation performed before it is stored in the scan converter
It is permanent and cannot be changed
  Done on a live image
Postprocessing is any manipulation of image data after storage. Occurs after A-to-D conversion. This can be reversed or restored to the original data.
  Always done on a frozen image
  (except some newer systems)
Notes: Except on some newer equipment

## Slide 32: Magnification
Also called “zoom” is the enlargement of a portion of an image to improve detail
  It fills the screen; the area
  magnified is called the
  region of interest (ROI)
There are two types of magnification with one of them more superior than the other
  Read magnification
  Write magnification

## Slide 33: Read Magnification
Read magnification
  Occurs after the image data is in a scan converter.
  Post-processing
  Similar to zooming in on a portion of a digital photo
Can be used on a frozen image

## Slide 34: Read Magnification
Occurs after image is stored; and has a 3-step process:
1. The sonographer scans an area and freezes the anatomical image.
2. Image is now converted from analog to digital and stored in the scan converter,
3. The sonographer determines the area of interest and utilizing system controls enlarges the image; the system reads and displays only the original data in the region of interest.

## Slide 35: Read Magnification

## Slide 36: Read Magnification
The new image has the same number of scan lines and pixels as the original image
Spatial resolution does not change because the pixel count is the same in the ROI
This process does enlarge each individual pixel.
This is performed on a frozen image – postprocessing.

## Slide 37: Write Magnification
Write Magnification
  occurs upon acquisition of the ultrasound reflections
  Preprocessing
  Just like retaking a digital picture using the zoom feature
Can improve temporal resolution (only if zoomed image is shallower than original image)
When the size of an image increases: if the pixel size is unchanged, the number of pixels increases

## Slide 38: Write Magnification
This is applied during data acquisition and requires a 4-step process:
1. The sonographer scans an area and freezes the anatomical image.
2. Image is now converted from analog to digital and stored in the scan converter.
3. The sonographer determines the region of interest. At that moment the system discards the old data in the scan converter
4. The system now reacquires an image only from the selected ROI and writes new data into the scan converter

## Slide 39: Write Magnification
Characterized by:
discarding of the old image when the new image data is acquired
the number of scan lines or pixels in the ROI is greater than they were in the same region of the original image
Spatial resolution improves in the ROI because the number of pixels have increased in the region enlarged. Pixel size is the same.
  Note the difference with write mag.: pixels were not just enlarged to make the image (read mag.)

## Slide 40: Coded Excitation
A sophisticated method of improving image quality
Coded excitation creates very long sound pulses containing a wide range of frequencies
Short pulses are used in traditional imaging (red)
Long complex pulses are used in coded excitation (green)

## Slide 41: Spatial Compounding
A method of using sonographic information from several different imaging angles to produce a single image. The frames are combined to form a single real time image.
This reduces speckle and minimizes shadow artifacts
It will reduce the frame rate and temporal resolution (Can only be performed with phased array transducers)

## Slide 44: Frequency Compounding
An advanced technique that reduces speckle artifact and noise in ultrasound images.
With frequency compounding, the reflected sound (black) is divided into parts with different frequency ranges, called sub-bands (blue, red, and green)

## Slide 45: Edge Enhancement
A method that makes the pictures look sharper.
The computer identifies and emphasizes sharp edge boundaries in the image.
It creates subtle bright and dark highlights on either side of these boundaries

## Slide 47: Temporal Compounding/ Persistence/ Frame Averaging
An imaging process technique that continues to display information from older images.
With persistence a number of prior frames are interposed onto the current frame.
A smoother image with reduced noise, higher signal to noise ration, and improved image quality is produce
This differs from spatial compounding in that all of the frames are from the same view/angle.

## Slide 50: Fill-in Interpolation

## Slide 51: Elastography
A dynamic technique that produces images from sound reflections in a new way.
Images are related to the mechanical properties of tissue. Depending on how tissues respond to force- they will be displayed in a different way.
It is used to help distinguish between benign and malignant masses (especially in the breast)

## Slide 53: PACS and DICOM
PACS- Picture Archiving and Communication System- the digital ultrasound storage lab (Computer Hard drives)
  Instant access
  No degradation of information
  Store and forward technology
DICOM- Digital Imaging and Computers in Medicine
  A set of rules and protocols that allow information to be stored, and then universally connected to a PACS System