## Slide 1: Quality Assurance
Chapter 22

## Slide 2: Quality Assurance
Routine periodic evaluation of the Ultrasound system to guarantee optimal image quality
Quality assurance is a medical and legal requirement for every laboratory.
Must be performed periodically and routinely

## Slide 3: Why?
We must validate image accuracy and quality.
Quality assurance is a method of testing our machine to make sure that everything is working as it should be and our images are accurate.

## Slide 4: Requirements
  Assessment of the system’s components
  Repairs
  Preventative maintenance
  Record keeping

## Slide 5: Goals
  Proper equipment operation
  Detection of gradual changes
  Minimize downtime
  Reduce number of non-diagnostic exams
  Reduce number of repeat scans

## Slide 6: Methods
Test under known, defined conditions
Constant instrument settings
Use phantom with measurable characteristics
  Image in identical environment
  A must for every ultrasound lab

## Slide 7: Devices
AIUM 100 mm test object
Tissue equivalent phantom
Doppler phantom
Beam profile/slice thickness phantom

## Slide 8: Quality Assurance
Quality assurance or QA is the responsibility of the SONOGRAPHER!!!

## Slide 9: AIUM 100mm test object
Fluid filled tank
Contains strategically located metal pins or plastic strings.
Speed of sound in the fluid is the same as that of soft tissue.
Does not have attenuation properties of soft tissue so it cannot evaluate gray scale
Commercially available

## Slide 10: Accuracy & Performance
Scanning a test object evaluates the accuracy and performance characteristics of a system.
Placing the transducer on the top, sides, and on the oblique side of the object provides a variety of orientations between the sound beam and the pins.

## Slide 11
Evaluate Dead Zone
Evaluates Lateral Resolution
Evaluates axial
Resolution
Electronic caliper accuracy (measurements) are evaluated by comparing the distances between reflections on the display with the actual distances in the test object. ​
Dead zone is evaluated by scanning the pins closest to the top of the object.

## Slide 12: AIUM 100 mm test object

## Slide 13: Tissue Equivalent Phantom
Commercially available
Strategically located pins, structures that mimic cysts and solid masses are embedded in the phantom
Sound speed in phantom is equal to speed in soft tissue
Attenuation coefficient similar to that of soft tissue
This phantom is useful for evaluating gray scale

## Slide 14: Tissue Equivalent Phantom

## Slide 15: Tissue equivalent phantom

## Slide 16: Doppler Phantom
Contains static simulated vessels as well as dynamic simulated vessels at a variety of angles
Pump forces echogenic fluid through “vessels” at known velocities. Constriction is found in one of the vessels as well.
Assesses accuracy of pulsed, continuous wave, and color flow systems
Vibrating sound and moving belt phantoms may also be used to evaluate Doppler systems

## Slide 17: Doppler Phantom

## Slide 18: Beam Profiler/Slice Thickness Phantom
  Tank with wires imbedded at increasing depths
  Used to display the characteristics of ultrasound beams as produced by the transducer
  Measures the beam geometry perpendicular to the imaging plane.
  Thinner slices improve spatial resolution, so we want to maintain the thinnest slice possible.
  If our slice is too thick, objects such as cysts may appear “filled in”

## Slide 19: Performance Measurements
Testing the different functions of the machine for accuracy
Several types of performance measurements

## Slide 20: Sensitivity
  Refers to the ability of a system to display low-level echoes.
  It is evaluated in a variety of ways.

## Slide 21: Normal Sensitivity
  Settings at which all pins, solid masses, and cystic structures in the test phantom are accurately displayed.
  Output power, TGC and amplification (gain) are adjusted to establish normal sensitivity.

## Slide 22: Minimum Sensitivity
  Assessment of the weakest echo signal that is accurately displayed.
  Determined by setting TGC flat then increasing system gain from the minimum value to the gain when a deep rod is displayed.
    The same rod in the test object should always be tested and settings should not vary from one evaluation to the next.

## Slide 23: Maximum Sensitivity
  Evaluated with the output power and amplifications set to maximum practical levels.
  Tissue phantom is imaged and depth of tissue like texture is determined.
  Maximum visualization depth is used to assess sensitivity.

## Slide 24: Dead Zone
  Region closest to the transducer where images are inaccurate.
  Results from the transducer ringing and the time it takes for the system to switch from transmit to receive mode
  Assessed with the most shallow series of pins in a test object.
  An increasingly deeper dead zone may indicate a cracked crystal, detached backing material, or longer pulse duration.

## Slide 25: Range/Registration Accuracy
  Describes the systems accuracy in placing reflectors at correct depths located parallel to the sound beam.
  If differences appear between ultrasound scan and actual pin locations in a test object the error may be caused by:
    System malfunction
    Speed of sound in phantom is less than 1540 m/s.

## Slide 26: Focal Zone
  Sonographer controlled phased array transducers must be carefully evaluated.
  Systems with dynamic receive focusing should produce narrow reflections over a wide range of depths.

## Slide 27: Axial & Lateral Resolution
Axial resolution is evaluated when the pins in the test object are parallel to the sound beam’s main axis.
Lateral resolution is evaluated when the pins in the test object are perpendicular to the sound beam

## Slide 28: Compensation Operation or Uniformity
  Describes the system’s ability to display similar reflectors in the phantom with echoes of equal brightness
  With proper TGC, identical reflectors should have the same appearance on the monitor, regardless of the depth

## Slide 29: Mock Cysts & Solid Masses
  Tissue equivalent phantom is used to evaluate dimension, texture, and fill-in of cysts.
  System should accurately display solid structures as hyperechoic and hollow structures as anechoic.

## Slide 30: Display, Hardcopy Output, and Gray Scale Dynamic Range
  Adjusting the systems output power and amplification should produce changes in the grayscale display
  It is important to compare the relationship between the image on the system’s screen with the output of all other display devices such as remote viewing stations (PACS), paper, film etc.
  How the image is displayed or sent for someone else to view must be very close to what is seen on your monitor on the machine!