## Slide 1: Chapter 14: Pulsed Echo Instrumentation

## Slide 2: The ultrasound system has two major functions
Preparation and transmission of electrical signals to the transducer, which creates the sound beam
Reception of electrical signals from the transducer with subsequent processing into clinically meaningful images and sounds

## Slide 3: Pulsed Echo Instrumentation
Six interconnected components – information is transferred to and from each component
1. Transducer
    Converts electrical energy into acoustic energy during transmission
    Converts returning acoustic into electrical energy during reception

## Slide 4: Pulsed Echo Instrumentation
2. Pulser/ beam former
    Component that controls the electrical signals sent to the transducer for sound pulse generation
    Determines PRF, PRP, and pulse amplitude
    Creates the firing pattern for phased array systems. This is called the Beam Former.

## Slide 5: Pulsed Echo Instrumentation
  3. Receiver
    Electronics associated with processing the electrical signal produced by the transducer during reception and producing a picture on an appropriate display
  4. Display
    Device associated with the presentation of processed data for interpretation

## Slide 6: Pulsed Echo Instrumentation
  5. Storage
    Devices and media that are used to permanently archive data
  6. Master Synchronizer
    Maintains and organizes timing and interaction of the system’s components (the boss)

## Slide 7: Pulser
Function:
  Receives timing signal from master synchronizer
  Creates electrical signal that will excite PZT crystal
  Produces electrical voltage 10-500 volts that excites PZT during transmission
    Greater voltage increases sound intensity
  The pulser controls transducer output

## Slide 8: Transducer Output (Power knob)
PZT vibrates with a strength related to the pulser voltage
When you change the transducer output, every pulse changes, image brightness changes
Higher voltage magnitude, higher transducer output, brighter image
Lower voltage magnitude, lower transducer output, darker image

## Slide 9: Transducer Output (Power knob) cont'd
Increasing transducer output improves signal to noise ratio.
  The meaningful signal strength is improved while the noise level remains constant.
What is noise?
  Random disturbance that obscures or reduces clarity of a signal.
  May contaminate images with low level, undesirable signals
With ALARA – first attempt to change your output power before changing overall gain if the image is too bright

## Slide 11: PULSE REPETITION PERIOD
  The pulser also determines the pulse repetition period (PRP)
  PRP determines the maximum imaging depth
Shallow Imaging | Deep Imaging
Shorter listening time | Longer listening time
Shorter PRP | Longer PRP
Higher PRF | Lower PRF

## Slide 13: Beam Former
Part of the transmitter
Functions with array transducers
Apodization – reduction of lobe artifacts
Modern systems- digital beam former
  Software programming
  Stable
  Versatile

## Slide 14
Switch protects the receiver from the extreme signals sent at transmission
Also directs the electrical signals to the appropriate processing component.
CHANNEL

## Slide 15: Receiver and its functions
The signals returning from the transducer are extremely weak
The receiver boosts the strength of these signals, processes them and prepares them for display.
Receiver
Display
The signals returning from the transducer are extremely weak​
The receiver boosts the strength of these signals, processes them and prepares them for display.

## Slide 16
ORDER OF OPERATIONS IN THE RECEIVER
  Amplification
  Compensation
  Compression
  Demodulation
  Rejection
You can remember the order that things will happen in because they are in alphabetical order

## Slide 17: Amplification (Overall gain)
Increases the strength of all electrical signals in the receiver prior to further processing
Synonyms – gain, receiver gain
Can be adjusted by the sonographer
Effect upon image
  every signal is treated identically
  amplification changes the brightness of the entire image.

## Slide 18: Gain
Increasing overall gain cannot create an image with uniform brightness.
Signal to noise ratio is unchanged when amplification levels change.
Preamplifier-may condition signal before it is amplified-often performed within the probe.

## Slide 19: Compensation (TGC)
Purpose: used to create image of uniform brightness from top to bottom.
Since attenuation is strongly related to path length, echoes returning from great depths have lower amplitudes than those returning from shallower depths.
  We must account for this in our image.
Compensation makes all echoes from similar reflectors appear identical regardless of their depth

## Slide 20: Compensation (TGC)
Synonyms: TGC (time gain compensation), depth compensation, swept gain
Adjusted by the sonographer
Compensation treats echoes differently depending on the depth at which they arise
Compensation makes an image equally bright at all depths

## Slide 21: TGC and frequency
Adjustments to TGC are related to transducer frequency:
  Higher frequency
    beam undergoes more attenuation
    More TGC must be used
  Lower frequency
    beam undergoes less attenuation
    Less TGC is needed

## Slide 22: TGC AND ATTENUATION
We have to account for the change in attenuation with the TGC controls.
  With higher frequency we must account for more attenuation so our TGC controls must be higher
  With lower frequency we must account for less attenuation so our TGC controls must be lower

## Slide 23: TGC curves
The TGC curve is the arrangement of the tiny knobs or switches in on the TGC
  Near Gain – small amount of compensation
  Delay – depth where variable compensation begins
  Slope – compensation corrects for increasing attenuation
  Knee – reflections are nearing maxunum compensation
  Far gain – maximum amount of compensation from the receiver

## Slide 24: TGC curves
The TGC curve is the arrangement of the tiny knobs or switches in on the TGC
TGC
Image Depth
Higher frequency curve
Lower frequency curve
Curve shifts up and to the right
Curve shifts downward and to the left
Less Amplification
More Amplification

## Slide 25: Key Terms
Near field – top of the image, area nearest the probe
Far field – bottom of the image, farthest from the probe
Near field
Far field

## Slide 26: EXAMPLE
Sample TGC Registry questions:
What will you do if you cannot see reflectors in the near field of your image? (or the far field)
A: adjust the TGC!!

## Slide 27: Compression
Alters the grayscale content of the image to one that the human eye can see
Reducing the total range of signals
  Smallest to largest signal
  How “black and white” an image will appear, versus a wider range of grays within that image

## Slide 28: Compression
Purpose: keeps signals within the operating range of the system’s electronics and the gray scale within the range of what the human eye can see
  Done without altering the relative relationships between signals: largest stays largest, smallest remains smallest
  Decreases the dynamic range of the signals
  Effect upon image: changes the gray scale mapping / characteristics of the image

## Slide 29: Demodulation
Changes the signal’s form to one more suitable for TV display
Cannot be adjusted by sonographer
Not seen on our images and we cannot adjust it. It is performed automatically by the system.
Has two steps:
    Rectification-turning all negative voltages into positive voltages. Corrects for or eliminates negative voltages.
    Smoothing or enveloping-puts a smooth line around signal “bumps” to even them out

## Slide 32: Rejection
Purpose: displays low level echoes only when clinically meaningful
Also known as suppression, threshold
Very low-level echoes may or may not be important. Reject determines whether they appear on an image.
Affects all low-level signals everywhere on the image but does not affect bright echoes.
Is adjusted by the sonographer

## Slide 34: Dynamic Frequency Tuning
Pulses contain a wide range of frequencies
On some us systems, higher frequencies create shallow parts of the image and lower frequencies create deeper parts
Echoes arising from superficial structures are filtered to process only higher frequencies since higher frequencies make better images
Lower frequency signals are used to image deeper structures since the higher frequencies have attenuated and are no longer present.
Takes advantage of the range of frequencies present for optimal imaging

## Slide 35: Output Power vs. Receiver Gain
Output power
  Affects brightness by adjusting the strength of the sound pulse sent into the body by the transducer
  When the pulse is more powerful, the signals returning are stronger and the image is brighter
  When the image is too bright, lateral and longitudinal resolution degrade
Receiver Gain
  Affects the brightness by changing amplification of the electronic signals after they return
  When amplification is increased, the electronic signals in the receiver are boosted and the image is brighter

## Slide 36: Output power vs. Receiver gain (cont'd)
Adjustments to either output power or receiver gain change the brightness of an entire image
Put simply, both can control brightness of the image
  Power controls intensity of transmitted sound
  Gain controls amplification of the signal at reception

## Slide 37: Output Power vs. Receiver Gain (cont'd)
How do we determine which to use?
On the registry:
  When a term in a question suggests outgoing, the function is probably output power (transmission)
  when a term suggests reception or incoming it likely refers to receiver gain (reception)

## Slide 38: Alara principle
When scanning:
  We use the ALARA principle to determine which to adjust.
    As Low As Reasonably Achievable
This means that we always choose the option that minimizes patient exposure while still obtaining a clinically relevant image
If your image is too bright, which should you choose?
    Decrease output power ---ALARA!!!!
If your image is too dark, which should you choose?
    Increase gain ---- ALARA!!!!

## Slide 41: What should you do?

## Slide 42: Turn down the power!

## Slide 43: What should you do?

## Slide 44: Turn down the gain!

## Slide 45: What should you do?

## Slide 46: ADJUST THE TGC! (MAKE THE IMAGE UNIFORM BRIGHTNESS FROM TOP TO BOTTOM)